PRESS [arrow keys] to scroll through this document:

Table of Contents:

1. Sokkia NET-2

2. Sokkisha SET-3

3. Cable Configurations

===================================================

SOKKIA NET2: (per manual received 07-01-94)

1. Mode selection:

a. Theodolite Mode {angles only}--> press [3]

b. Basic Mode {angles + EDM}:

Polar Distance (slope dist)-->   press [CE-CA][7]

2. Precision Plumbing using instrument tilt angles:

a. press [ENT][0]

b. press [3]

c. wait a few sec; press [ENT][0]

d. rotate 180:

e. press [ENT][0]

f. level to make tilt <'s go to zero.

g. press [3] or [CE-CA]

3. Instrument Parameters from Inst. Keyboard:

a. press [CE-CA][1]

b. press [2] to Config.

c. press blue scroll keys up & down to pick/view various parameters.

To change a desired parameter:

press [ENT]

press [#]

press [CE-CA] to retain previous selection

4. Communication settings required for ET:

a. 1200 baud, Checksum, No Parity Bit

To Exit: PRESS [Alt],[ENTER],[X]

===================================================

===================================================

SOKKISHA SET-3: (from Lietz/Sokkisha Operator's Manual)

1. Specifications:

a. Angle Accuracy= +-5 arc-sec  (1 arc-sec minimum display)

b. EDM Accuracy= +-(5mm + [3ppm * D]) = ~+-.200 inch for 10 to 100ft.

c. Minimum Focus= 4.3 ft.

d. External Battery #6661-02 = 7 Hr. =~ 4000 measurements (Dist & Angle).

Internal Battery #6651-01 = 13 Hr. (Angle) =~ 600 meas (Dist & Angle).

2. Set Internal Switches:

a. Turn Power OFF (toggle switch on instrument side).

b. Remove cover on side of instrument standard, slide switches per

below:

7 __* 5"     {Angle Resolution 1"

6 __* Index  {Auto Index V Circle via plunging telescope

O  5 __* Tilt   {V Circle Compensator is ON

N  4 *__ Feet   {Display Distance In Feet

3 __* C&R    {No Correction for earth curvature or refraction

2 __* 0+-90  {V angle orientation controlled by switch 1

1 __* 0      {0 angle vertically up (Zenith) @ top of circle

3. Instrument Calibration (Optional):

a. Check Tilt Angle Compensator:

1. Inst. ON, V circle indexed.

2a. PRESS [ 0 SET ]

2b. PRESS [0 number] {This displays tilt angle.

3. Record tilt angle (a). {upper window, include +or- sign on left.

4. Rotate standard H 180 arc-deg +- 5 arc-sec.

5. Record tilt angle (b). {upper window, include +or- sign on left.

6. Calculate c.  c = (a + b)/2 . If c > 5 arc-sec then

a. Remove sensor index cover (R side of standard below battery).

b. Rotate back to 0 H.

c. adjust flat-head screw so upper tilt display = a - c.

d. Rotate 180 H.

e. adjust screw until upper display = b - c.

b. Reticle:

1. Select a good target @ 160 to 330 ft.

2. Inst ON, V Indexed.

3. Sight target (Direct) and record H1 & V1.

4. Sight target (In-Direct) and record H2 & V2.

5. Calculate Dh= H2 - H1 and Dv= V2 - V1.

6. If { Dh <> 180 +- 20" } or { Dv <> 360 +- 20" } then

adjust:

a. Remove two screws holding the reticle adjustment cover (above

and below eyepiece).

b. In the In-Direct position, using tangent screws, adjust:

H to read:  (H1 + H2)/2 + 90

V to read:  (V1 + V2)/2 + 180

c. move reticle using the adjusting screws to center the reticle

on the target.

7. Repeat steps 3 thru 6 until within tolerance.

c. Check EDM:

1. Check EDM on line against a known reference (assure variable

parameters are correctly set per below).

4. Set Instrument Parameters:

a. Precision Plumb Instrument.

b. Select Mode:

1. ET, Theodolite Mode (V & H Angle Output {no distance}):

a. Turn Power ON (toggle switch on instrument side).

Observe display:

{ V   0   _|_+ }

{ H >   00000  }  note : the > indicates horizontal angle right.

PRESS [CE] to scroll thru options.

NOTE: ET must use angle right.

b. Plunge telescope to index V circle.

c. PRESS [top right large button which looks

like a transit]

d. Optional: PRESS [0 SET] to zero H angle.

2. ETTS, Theodolite Mode + EDM (V, H, & Distance):

a. Set PPM adjustment on EDM {example: [+][1][3]

1. Obtain the Temperature (+-2F) & Atmospheric Pressure (+-.2InHg

or +- 5MMHg).

2. Use the Atmospheric Correction Chart to determine PPM.

b. Set the [MEAS./TRACK] swith on the EDM to MEAS.

c. Set Prism Constant switch (3 {cm} for typical prisms) located

behind ~.75in dia screw on telescope front side.

d. Turn Power ON (toggle swith on instrument side).

e. Plunge telescope to index V circle.

f. PRESS [CA],[7],[CA] keys. This places the instrument into EDM

mode and allows ETTS to poll the instrument.

Note: The instrument must be manually placed in EDM mode after

using theodolite mode (say for setting-out {StakeOut}) in order

for ETTS to successfully communicate with the instrument.

The SET-# instruments (no sub letter) were pre 2-way

communications thus it is impossible for ETTS to tell

the instrument to switch back to basic EDM mode.

g. Notes: PRESS [CA] to stop measurement.

PRESS [(|)] to power ON or OFF EDM.

c. Set Remaining Parameters:

a. The SET3 always transmits in Degrees+Minutes+Seconds.

The EDM, per above, is set to measure in Feet.

If you have selected other units in a previous ET

menu then you must go back and correct.

d. Error Messages:

[ b. dE d ]  = Battery Voltage is too low.

[ E  100  ]  = Error when measuring a H angle.

[ E  101  ]  = Error when measuring a Zenith angle.

[ E  115  ]  = Exceeding Compensator angle range (+-3').

[ E  117  ]  = Exceeding Compensator angle range (+-3').

[ E  200  ]  = Incoming light problem (EDM).

[ E  201  ]  = No reflection light for EDM.

[ E  208  ]  = Exceeding distance +-19999.99

To Exit: PRESS [Alt],[ENTER],[X]

===================================================

===================================================

Typical RS232 Cable Configurations (Ref. Only):

1:

Inst          Inst Cord          Std RS232 Cable      Computer

[SET-#] [Limo Conn]-----[DB-9] [DB-9]----------[DB-9] [PC]

2:

Inst          Inst Cord                              Long-Range Cable                                  Std RS232 Cable   Computer

[SET-#] [Limo Conn]-----[DB-9] [3PMF power] [DB-25/1000$/DTE]--[DB-25/1000$/DCE] [3PMF power] [MP Box] [DB-25]----[DB-9] [PC]

^--+V                                             ^--+V        ^--+V

3:

Inst          Inst Cord          Std RS232 Cable                  Std RS232 Cable     Computer

[SET-#] [Limo Conn]-----[DB-9] [DB-9]----------[DB-25] [MP Box] [DB-25]--------[DB-9] [PC]

^--+V

Sokkia PigTail Cord Circuit (08-17-95 using SET-3 {#81846):

Pin #        Pin #

LIMO CONN    DB-9 CONN

(3)----------(2)

(4)----------(3)

(1)----------(5)

Cable Power test @ 1009$ Conn (ref only):

note: current draw of 1009$ is 10mw @ 9V.

+9v dc ------(pin 4/DB-9)

-9v dc ------(pin 5/DB-9)

===================================================

TO EXIT:

If using MS-DOS Editor--> PRESS [Alt],[ENTER],[X]

For Other Editors--> Follow specific editors menus.